## Slide 1: Development of Computer Vision and Image Processing Libraries at NSLS-II
William Watson
Kazimierz Gofron

## Slide 2: Purpose
Development of image analysis software integrated with EPICS control system.
Computer Vision for synchrotron beamline applications.
Develop easy to use python modules to access computer vision functions backed by OpenCV.
Optimized results for fast computation, via C++, Intel IPP/TBB.
Automate processes from image input such as:
  Positon, spread, and intensity of X-Ray Beams
  Isolating and computing data on multiple objects
  Provide assistance to mounting samples and report any errors

## Slide 3: BPM-1 X-Ray Analysis (IXS)
Images Courtesy of: Yong Cai, Alessandro Cunsolo, Alexey Suvorov, and the IXS Beamline Staff
Purpose:
Discover Position of Beam
Discover Spread of Beam
Discover Intensity of Beam
Discover the Centroid of the Beam
E = 9.1 keV
Scintillator: YAG
Camera: Prosilica
Magnification: 5x

## Slide 4: BPM-1 X-Ray
3rd XBPM
2nd XBPM
1st XBPM
Images Courtesy of: Yong Cai, Alessandro Cunsolo, Alexey Suvorov, and the IXS Beamline Staff
Located in B Hutch (IXS)

## Slide 5: BPM-1 Data Results
Console Output:
Object Details:
perimeter: 2356.99022925
orientation: 179.838363647
max: (925, 198)
height: 372
extrema: {'B': (938, 568),
'R': (1054, 415),
'L': (813, 377),
'T': (914, 196)}
area: 65058.5
min: (1047, 564)
sum intensity: 20426526
width: 241
centroid: (933, 382)
mean intensity: 227.842390577
Images Courtesy of: Yong Cai, Alessandro Cunsolo, Alexey Suvorov, and the IXS Beamline Staff
Result Superimposed on Original
Thresholding Result
Jet Color Map

## Slide 6: Analysis -Merlin Quad X-Ray Detector (IXS)
Images Courtesy of: Yong Cai, Alessandro Cunsolo, Alexey Suvorov, and the IXS Beamline Staff
Data Retrieved from the Merlin Detector at IXS
Purpose:
Isolate Individual Streaks
Process each Streak to learn Position, Center, Spread, Intensity, Max Values, etc.
Count Intensity in each streak
E = 9.1 keV
Direct detection
PEL size = 55 [um]

## Slide 7: Merlin Data Results for First Object (Largest)
Console Output:
Object 1:
perimeter: 125.840619564
orientation: 179.981033325
max: (131, 78)
height: 55
extrema: {'B': (129, 122), 'R': (135, 98),
'L': (126, 92), 'T': (132, 67)}
area: 270.5
min: (134, 83)
sum intensity: 62689
width: 9
centroid: (130, 95)
mean intensity: 126.644444444
Images Courtesy of: Yong Cai, Alessandro Cunsolo, Alexey Suvorov, and the IXS Beamline Staff

## Slide 8: ABBIX Beamlines (AMX) – Pins in a Robotic Gripper
Purpose:
Find if Pin is in the Gripper
Determine if Pin is properly mounted
  Analyze the position of the Gripper and Pin within the image
  Find the Region of Interest (ROI) for the Pin and Gripper
  Compare the Pin and Gripper to a ‘Perfect’ Image to determine if anomalies are present
  Advise the user to any potential problems
  Discover kinks, check mounting
  Discover center of mass
  Use extrema for alignment assistance
Images Courtesy of: Jean Jakoncic, AMX Beamline

## Slide 9: ABBIX Beamlines (AMX): Pins in a Robot Gripper
Pin and Gripper ‘Perfect’ Image
Pin Template Image
Gripper Template Image
Images Courtesy of: Jean Jakoncic, AMX Beamline
Images are from AMX, Displays a Pin and a Gripper, in an attempt to grab the pin

## Slide 10: Pin/Gripper – ‘Perfect’ vs Test Cases
‘Perfect’
Image 1
Image 2
Image 3
Image 4
Images Courtesy of: Jean Jakoncic, AMX Beamline

## Slide 11: Pin/Gripper Image Analysis – Results
Images Courtesy of: Jean Jakoncic, AMX Beamline

## Slide 12: Crystal Rotational Alignment - AMX
Purpose:
Locate the Crystal
Center the Crystal in Goniostat
Plot as a function of angles the Y Pixel coordinate of Crystal during rotation
Images Courtesy of: Jean Jakoncic, AMX Beamline

## Slide 13: Crystal Rotational Alignment - AMX
Centering Crystal:
Images Courtesy of: Jean Jakoncic, AMX Beamline

## Slide 14: Loop Centering - AMX
Images Courtesy of: Jean Jakoncic, AMX Beamline

## Slide 15: Conclusion
Computer Vision is integrated with EPICS control system.
Computer Vision provides:
  Center Samples in Goniostat (ABBIX, IXS, etc.)
  Automated Robotic Mounting and Sample Detection (AMX)
  Assist in alignments of samples, crystals, and beams
  Discovers and reports information about objects within an image
  Prevent potential problems by alerting users of anomalies
With OpenCV 3.1.0 and the Intel IPP Library, all CV functions run at optimized speeds, providing the best computational results currently available.